## Title

## An evaluation of Qiagen HPV sign for the detection and genotyping of cervical lesions and oropharyngeal squamous cell carcinomas

Authors

Chris Warda,d, Johanna Pedrazaa, Kimberley Kavanaghb, Ingolfur Johannessenc, Kate Cuschieria.

aScottish Human Papillomavirus Reference Laboratory, Directorate of Laboratory Medicine, Royal Infirmary of Edinburgh, 51 Little France Crescent, Edinburgh, EH16 4SA, United Kingdom

bDepartment of Mathematics & Statistics, University of Strathclyde, Livingstone Tower, 26 Richmond Street, Glasgow, G1 1XH, United Kingdom

cEdinburgh Specialist Virology Centre, Directorate of Laboratory Medicine, Royal Infirmary of Edinburgh, 51 Little France Crescent, Edinburgh, EH16 4SA, United Kingdom

dCorresponding Author:  Chris.Ward@nhslothian.scot.nhs.uk, Tel: +44 (0)131 242 6010

## Abstract

HPV genotyping is an important tool in the epidemiology and surveillance of HPV associated cancers and for the risk-stratification of HPV infections.  HPV sign Genotyping Test (QIAGEN) is a new pyrosequencing assay for the detection and genotyping of HPV. The sensitivity and comparative performance of HPV sign was determined using a sample panel derived from histologically confirmed cervical lesions (cervical intraepithelial neoplasia 2 or worse) and oropharyngeal squamous cell carcinomas. Comparative analysis showed that 80% of cervical intraepithelial neoplasia 2+ and 81% of  oropharyngeal squamous cell carcinomas  were HPV positive by HPV sign compared to 100% of the cervical intraepithelial neoplasia 2+ and 81% of  oropharyngeal squamous cell carcinomas   by the digene HPV Genotyping RH Test (RH), and INNO-LiPA HPV Genotyping Extra assay respectively. Fewer genotypes were detected overall by HPV sign than via the relevant comparator assays (10 v 21 for cervical intraepithelial neoplasia 2+; 4 v 9 for  oropharyngeal squamous cell carcinomas) and also fewer multiple infections (9 v 28 for cervical intraepithelial neoplasia 2+; 0 v 4 for  oropharyngeal squamous cell carcinomas). HPV sign results were more compatible with the comparator assay for the  oropharyngeal squamous cell carcinoma samples (100%) than for the cervical samples (73%). These results suggest that HPV sign in its current form is  suited to samples that harbour multiple infection.

Key Words

HPV, pyrosequencing, genotyping, cervical, oropharyngeal squamous cell carcinoma

1. Introduction

Persistent infection with a high-risk HPV type is necessary for the development of certain anogenital cancers including cervical cancer (Walboomers  et al. 1999), HPV infection is associated also with an increasing proportion of Oropharyngeal Squamous Cell Cancers (van Monsjou et al. 2010) with HPV positivity indicating an improved prognosis. Currently 12 HPV types have been characterised as class I carcinogens for humans, and several other HPV types classed as possibly carcinogenic to humans, by the International Agency for Research on Cancer (IARC 2012).  HPV genotyping is vital for epidemiology and surveillance of HPV, particularly given the introduction of immunisation programmes. Genotyping may also have an increasing role in the risk stratification of infection for HPV associated disease management (Khan et al 2005).

Currently there are more than 100 commercially available HPV assays (Cuschieri 2011, Poljak et al. 2012) with output results ranging from limited typing of HPV 16 and 18 (e.g. cobas® HPV Test (Roche) and the Real Time High Risk HPV Test (Abbott) through to middle-range typing of 12 to ~35 types via reverse line blot assays to multiplex detection of >80 types such as the Origene Luminex Multiplex system (Schmitt et al. 2011).  Pyrosequencing technology offers a new approach to HPV detection that potentially broadens the ability to detect rare or novel virus genotypes (Barzon et al. 2011).

HPV sign Genotyping Test (HPV sign) (QIAGEN, Hilden, Germany) adapts pyrosequencing technology to HPV diagnostics and genotyping, however data on clinical performance is limited : Barbieri et al. (2012) in an analysis of 87 archived DNA extracts from cervical biopsies (34/87) and cervical swabs (53/87) found HPV sign had a lower sensitivity (61% v 69%) and ability to detect multiple infection compared to historical INNO-LiPA HPV Genotyping Extra assay (INNO-LiPA) (Innogenetics NV, Ghent, Belgium) data; one further study used HPV sign to screen cervical samples from attendees at a colposcopy unit, detecting HPV DNA in 50.5% in these patients (Paba et al. 2012).

A key performance attribute of any clinically orientated HPV assay is sensitivity for the detection of high-grade cervical lesions. Key performance attributes of an assay orientated to epidemiology and surveillance are an ability to delineate multiple infections and amenability to diverse bio-specimens.  Thus, the objective of this study was to consolidate and enhance the performance data on HPV sign using an enriched panel of previously annotated samples associated with cervical intraepithelial neoplasia 2+ and oropharyngeal squamous cell carcinomas.

2. Study Design

2.1  Sample Selection

The present study utilised archived, previously annotated nucleic acid (NA) collected as a result of either National HPV Immunisation Surveillance or routine/referral and audit associated with the services provided by the Scottish HPV Reference Laboratory. In total 100 clinical extracts collected within Scotland were assessed by HPV sign. Samples come under the NHS Lothian ‘Safe Haven’ for tissue research (NHS Lothian SAHSC Human BioResource:08/S1101/41 and 10/S1402/33)

The cervical cases consisted of NA derived from 47 cases of cervical intraepithelial neoplasia 2 and 32 cases of cervical intraepithelial neoplasia 3 (i.e. 79 cases of cervical intraepithelial neoplasia 2+) diagnosed in Scotland in 2005 and stored as formalin fixed paraffin embedded (FFPE) blocks as per routine practice.  10 micron sections were taken from the original FFPE block and stored at room temperature prior to extraction.  NA was extracted in 2009 using an adaptation of the QIAamp DNA Mini Kit (QIAGEN) optimised for FFPE sections (Steinau 2011) and all cases were genotyped using the digene HPV Genotyping RH Test (QIAGEN). Residual NA had been stored at -80°C and subject to only 1 freeze thaw cycle before the present analysis.

In addition a total of 21 NA extracts derived from  oropharyngeal squamous cell carcinoma samples were assessed. Cases had been received and extracted over a period encompassing 2009-11 with the extraction methodology as above. All  oropharyngeal squamous cell carcinomas  had been tested previously using the INNO-LiPA assay. Residual NA had been stored at -800C and had not been subject to any freeze thaw cycle before the present analysis.

2.2  HPV Genotyping

The HPV sign Genotyping Test (QIAGEN) was performed according to the manufacturers’ instructions. HPV sign uses EVA Green real-time PCR with primers for the hypervariable region of the HPV L1 ORF to amplify DNA from a broad range of HPV types using the Rotor Gene Q. Primers against human beta-globin are included to provide a control for cellularity/inhibition. The presence or absence of HPV DNA is determined by melting curve analysis. Samples determined to be HPV positive on the melt curve are advanced to pyrosequencing using the Pyromark Q24 with 4 specific sequencing primers recognising HPV 31-like, HPV 16-like, HPV 39-like, and a broad spectrum of HPV types, for the generation of sequence data. HPV types are identified from the sequence data using the IdentiFire version 1.0.5.0 software package (Biotage AB, Uppsala, Sweden).

Samples which tested as “no template” (NT) with HPV sign, or which had HPV sign genotyping results incompatible with the archived data from comparator assays, were retested with HPV sign at a 1/10 dilution of template to mitigate the effect of pcr inhibitors potentially present in the sample.

The research use only (RUO) RH (QIAGEN) incorporates a polymerase chain reaction stage using GP5+/6+ primers to amplify a region of the HPV L1 ORF. Amplicons are then hybridized to strips with 18 immobilized probes corresponding to all currently known high-risk HPV genotypes and probable high-risk HPV genotypes (16, 18, 26, 31, 33, 35, 39, 45, 51, 52, 53, 56, 58, 59, 66, 68, 73, 82). This version of the assay does not contain an internal control for sample adequacy.

INNO-LiPA uses SPF10 modified primers to PCR amplify a 65bp region of the HPV L1 ORF with human HLA-DPB1 primers used as an internal control for template quality. Amplicons are used as probe in a reverse line blot hybridization against DNA sequences from 18 high-risk and probable-high risk HPV genotypes listed above and an additional 10 HPV genotypes (6, 11, 40, 43, 44, 54, 69, 70, 71,74).

2.3  Data Analysis

A simple HPV positivity rate was determined for each assay and sample type, with cervical results stratified into cervical intraepithelial neoplasia 2 and cervical intraepithelial neoplasia 3. Concordance of genotyping results was assessed for HPV sign v. RH (cervical intraepithelial neoplasia 2+), and HPV sign v. INNO-LiPA (oropharyngeal squamous cell carcinoma  samples). Individual genotyping results were considered to be (i) fully concordant where an identical list of HPV genotypes was obtained by each assay; (ii) partially concordant where at least one genotype in common was assigned by each assay but where full concordance did not occur; and (iii). non-concordant (incompatible) where either the sample tested as positive by one assay but negative by another, or where there was no overlap in the genotypes assigned by each assay.

Assay specific counts were taken for the occurrence of specific HPV genotypes within the cervical and  oropharyngeal squamous cell carcinoma sample sets. HPV detection was stratified into mono v multiple infections for : (i) all positive samples; (ii) all high-risk HPV types (see above); (iii) specific high-risk HPV genotypes of particular interest to immunisation surveillance (HPV 16, 18, 31, 33 and 45). Kappa values (Cohen 1960) and associated 95% confidence intervals were calculated as a measurement of agreement between HPV sign v RH (cervical samples), and HPV sign v INNO-LiPA (oropharyngeal squamous cell carcinomas).

3. Results

3.1  Cervical Samples

3.1.1 Overall HPV Positivity of cervical intraepithelial neoplasia 2+ by HPV sign

Of the 79 cervical biopsies associated with cervical intraepithelial neoplasia 2+, the Rotor Gene Q identified 61 samples as HPV positive (77.2%), and these were advanced to pyrosequencing stage as per the manufacturer’s algorithm.  It was notable that only one sample gave a positive result for the template control (human beta-globin) in the Rotor Gene Q analysis.

A positive genotyping result was obtained from 59/61 pyrosequencing reactions, giving an overall HPV positive rate of 75% (95% CI: 64%, 83%) by HPV sign in the cervical intraepithelial neoplasia 2+ panel. Follow up HPV sign testing of initially HPV negative samples (n = 20) at 1/10 dilution resulted in a further four positive genotyping results, increasing the HPV positivity rate to 79.7% (95% CI: 70%, 87%). The HPV positive rates were similar for cervical intraepithelial neoplasia 2 (80.9% CI 67.5%, 89.6% n= 47) and cervical intraepithelial neoplasia 3 (78.1% CI 61.2%, 88.9%  n= 32).  All of these samples had tested previously as HPV positive by RH.

3.1.2 Type specific concordance RH vs HPV sign in cervical intraepithelial neoplasia 2+

There was full concordance between the RH and HPV sign genotyping results for 38% (95% CI: 28%, 49%) of the cervical samples with a further 35% (95% CI: 26%, 46%) of results displaying partial concordance i.e. at least one HPV type detected in common. For 21 samples (26%, 95% CI: 18%, 37%) the RH and HPV sign results were incompatible. In 16 of these cases the sample tested as HPV positive by RH, but HPV negative by HPV sign, this including one instance of HPV 16, one instance of co-infection with HPV 18 and HPV 51, and five instances of HPV 31.  The  5 remaining cases tested as HPV positive by both RH and HPV Sign, however  non-overlapping  sets of genotypes were identified by RH and HPV Sign (Table 1).

The RH results describe a more complex population of HPV types (21 HPV types, 79 positive results, 28 multi-type infections with up to 4 co-infection types) compared to the HPV sign results (10 HPV types, 63 positive results, 9 mixed infections with a maximum of two types within a sample (Tables 2 & 3).

The genotypes identified most frequently by RH were HPV 16, 18, 33 and 31. HPV sign detected similar numbers of HPV 16, 18 and 33, but notably fewer cases of HPV 31 (13 cases v 2 cases for RH and HPV sign respectively). Table 2 shows the occurrence of selected high-risk HPV types in both mono- and multiple infections as measured by each assay.  Overall, 73/79 (92%) cervical samples were positive for high-risk HPV types (HPV 16, 18, 31, 33, 35, 39, 45, 51, 56, 58, 59 and 68) by RH, compared to 57/79  (72%) samples by HPV sign (Table 2).

35% of samples testing positive by RH contained >1 HPV genotype, including 5 samples which contained 4 co-infecting HPV types. However only 14% of HPV sign positives contained multiple types and no more than 2 HPV types were identified from a single sample. Kappa values for HPV sign v RH are considered substantial  (Landis JR, Koch 1977)  for HPV 16, 18, 31 and 33, but only fair for HPV 45, and poor for “all high-risk HPV” (Table 2). A full listing of genotypes obtained from individual cervical samples is provided in supplementary table S1.

3.2   Oropharyngeal squamous cell carcinomas  samples

3.2.1 Overall HPV positivity of  oropharyngeal squamous cell carcinomas  by HPV sign

Twenty-one  oropharyngeal squamous cell carcinoma  samples tested previously by INNO-LiPA HPV Genotyping Extra Test were selected for analysis by HPV sign. Seventeen (81%) had tested HPV positive by INNO-LiPA. HPV sign identified 15 cases of HPV, rising to 17 (81% HPV positivity) following reanalysis of initially HPV negative samples at 1/10 dilution. The same subset of  oropharyngeal squamous cell carcinomas samples tested as HPV positive by INNO-LiPA and HPV sign.

3.2.2 Type specific concordance RH vs HPV sign in oropharyngeal squamous cell carcinomas

RH and HPV sign genotyping results for  oropharyngeal squamous cell carcinoma  samples were 100% compatible following two rounds of HPV sign testing (neat sample and 1/10 dilution), with full agreement achieved in 81% (19/21) of samples. Restricting analysis to the HPV sign performed on neat samples yields 2 discordant results (9.5%), both of which were HPV 16 positive by INNO-LiPA but HPV negative by HPV sign (Table 1).

Fewer HPV types were identified in the  oropharyngeal squamous cell carcinoma samples  than in the cervical samples, with a total of 9 HPV types identified by INNO-LiPA, and 4 by HPV sign (Table 3). As expected, HPV 16 dominated results from both assays, being present in 57% of samples by both INNO-LiPA and HPV sign.

None of the  oropharyngeal squamous cell carcinoma samples contained more than a single genotype according to HPV sign, however the RH assay identified 4 incidences of multiple co-infecting HPV types. The small number of oropharyngeal squamous cell carcinoma  samples and the limited heterogeneity of types, limits the use of the Kappa statistic, particularly for individual HPV types.  A full description of the individual genotypes obtained from  oropharyngeal squamous cell carcinoma  samples is given in supplementary table S2.

Internal template control results were consistently negative for all  oropharyngeal squamous cell carcinoma  samples for both INNO-LiPA and HPV sign whether tested as neat sample or at a 1/10 dilution.

4. Discussion

In this study, a sample set of DNA extracted from 79 cervical intraepithelial neoplasia 2+ biopsies which had tested previously as HPV positive by the RH assay were tested using the new HPV sign genotyping assay. HPV sign performed on undiluted samples detected and assigned an HPV genotype in 59/79 (75%) of cases, rising to 63/79 (80%) upon re-analysis of initially negative samples at 1/10 dilution. There was no significant difference in the detection rates for cervical intraepithelial neoplasia 2 and cervical intraepithelial neoplasia 3 samples using HPV sign. For an HPV assay to be clinically applicable - a sensitivity of 90% for cervical intraepithelial neoplasia 2+ (relative to a gold standard assay) is suggested and according to these data HPV sign falls just below this threshold.

A smaller range of HPV types and fewer multi-type infections were detected by HPV sign compared to RH in the cervical intraepithelial neoplasia 2+ samples. Further, HPV sign detected a maximum of two co-infecting HPV types, compared to four with RH, and never more than one type from within one of the four pyrosequencing channels used in the Pyromark Q. Certain HPV types were identified less often by HPV sign than by RH, notably the high-risk types HPV 31, 52 and 56 (Table 3), Approximately one-third of cervical genotyping results were in full concordance between HPV sign and RH, with a further third of samples giving compatible (partially concordant)  results. The study population did not incorporate any cervical intraepithelial neoplasia 2+ that tested negative for HPV by RH. Although, HPV negative cervical intraepithelial neoplasia 2+ (particularly cervical intraepithelial neoplasia 3+) are rare, it is possible that concordance between HPV sign and RH results may be higher in a sample which includes HPV negative biopsies; consequently HPV sign may be disadvantaged in this respect.

For the  oropharyngeal squamous cell carcinoma  samples 100% concordance between RH and HPV sign results was reached at the qualitative level when neat and diluted sample results were aggregated for the latter. The initial round of HPV sign testing on neat samples yielded an HPV detection rate of 88% for the  oropharyngeal squamous cell carcinoma    samples, this is higher than the 74.7% observed with the cervical intraepithelial neoplasia 2+ samples.

As in the cervical samples, HPV sign identified a narrower range of HPV types in the  oropharyngeal squamous cell carcinoma samples than did the comparator line blot assay. None of the  oropharyngeal squamous cell carcinoma samples contained more than one HPV genotype according to HPV sign. Concordance and kappa scores for HPV sign v INNO-LiPA were much higher than for HPV sign v RH in the cervical samples, indicating that in its current form HPV sign may be suited more to sample types that are less likely to contain multiple infections.

The internal control for sample adequacy in both INNO-LiPA (no positives from twenty-one samples) and HPV sign (one positive result from seventy-nine samples) performed poorly in this study, which is in contrast to the internal control for the Luminex HPV assay which has been shown to perform well with FFPE material (Schmitt et al. 2011).   This is a particular concern in the clinical setting as the absence of a reliable cellular control makes it impossible to distinguish a true negative result from an inadequate reaction.  As described in the methods section, nucleic acid extraction was performed using a technique optimised for FFPE material and a kit that is provided by the same manufacturers of the HPV sign assay (Qiagen). Consequently, we would suggest that extraction technique is unlikely to have had a significant impact on internal control performance.  Re-consideration of the detection cut-for the existing internal control or reselection of an alterative target sequence may be warranted.

There are limitations to this study; the absolute number of samples is relatively small, although, all samples were either cancerous (oropharyngeal squamous cell carcinomas), or high-grade lesions (cervical intraepithelial neoplasia 2+),  .However, enriching for high grade lesions at least allows for a  robust assessment of sensitivity - a key attribute of any clinically orientated HPV test. Furthermore this study was not a true head to head; sampling, extraction and testing of the different assays was not performed at the same time and in this way the HPV sign was disadvantaged.  However, DNA stored at -800C has been shown to be stable for HPV detection, in a study where DNA had been stored for 8 years at -80 and subject to 5 freeze thaw cycles, high recovery and concordance for type specific HPV detection was observed (95%) ( Cuschieri 2008).

To conclude, the data suggest that further calibration of HPV sign to disease endpoints may be justified if it is to be considered for use in cervical disease management. In terms of its use for epidemiological work, the diversity and extent of multiple infection was lower compared to the comparator line blot assays. The higher concordance of HPV sign for the  oropharyngeal squamous cell carcinoma cases (compared to cervix) which harbour fewer types (and diversity) indicates that it may be more suitable for samples that contain fewer infections, at least in its current form.  One possible advantage of HPV sign that was not addressed in this study is its potential to detect rarer and/or cutaneous HPV types which are not incorporated into other HPV commercial assays. This latter feature may make HPV sign a good choice for investigative work in less established disease contexts and further work in this area would be of value.

5. Acknowledgements

We thank QIAGEN for the supply of HPV sign kits for this study.

6. Conflict of Interests

Qiagen provided the consumables and equipment required to perform HPV SIGN gratis, however, analysis and interpretation of the data was performed independently

7. References

Barbieri D, M Nocera, G Gallinella et al . Comparison of HPV sign Genotyping Test with INNO–LiPA HPV Genotyping Extra assay on histologic and cytologic cervical specimens. Diagn Micr  Infect Dis 2012;74,43–8

Barzon L, E Lavezzo, V Militello, et al.  Applications of Next–generation Sequencing Technologies to Diagnostic Virology. Int Journal of Mol Sci 2011,12: 7861–84

Cohen J. A coefficient of agreement for nominal scales. Educ Psychol Meas 1960;20, 37–46

Cuschieri K.  Monitoring trends in HPV diversity - Making the most of existing archives. Presented at 8th International Multidisciplinary Congress of the European Organisation for Genital Infections and Neoplasia, November 2008, Nice (ss 5-7)

Cuschieri K . Human papillomavirus testing: the challenges of picking the right tools for the job. Expert Rev Obstet Gynecol 2011;6:1 – 11

IARC Working Group on the Evaluation of Carcinogenic Risks to Humans. A review of human carcinogens. IARC Monogr Eval Carcinog Risks Hum. 2012;100(Pt B):1-441

Khan MJ, Castle PE, Lorincz AT et al.  The elevated 10-year risk of cervical precancer and cancer in women with human papillomavirus (HPV) type 16 or 18 and the possible utility of type-specific HPV testing in clinical practice. J Natl Cancer Inst  2005; 97: 1072-9

Landis JR, Koch GG. The measurement of observer agreement for categorical data. Biometrics 1977;33, 159–74

Meijer CJ, Berkhof J, Castle PE et al. Guidelines for human papillomavirus DNA test requirements for primary cervical cancer screening in women 30 years and older. Int J Cancer 2009; 124: 516-520

van Monsjou HS, Balm AJ, van den Brekel MM, Wreesmann VB . Oropharyngeal squamous cell carcinoma: a unique disease on the rise?  Oral Oncol 2010;46,780-5

Paba P, Ascone C, Criscuolo AA et al.  Human papillomavirus molecular testing in women with low grade cervical lesions: experience from an Italian hospital. Anticancer Res 2012;32,1253–7

Poljak M, Cuzick J, Kocjan BJ et al.. Nucleic acid tests for the detection of alpha human papillomaviruses. Vaccine. 2012;20;30 Suppl 5:F100-6

Schmitt M, de Koning MN, Eekhof JA et al.. Evaluation of a novel multiplex human papillomavirus (HPV) genotyping assay for HPV types in skin warts. J Clin Microbiol 2011,49:3262–7

Steinau M, S Patel and E Unger. Efficient DNA Extraction for HPV Genotyping in Formalin–Fixed, Paraffin–Embedded Tissues. JMD;2011:13 ,377 – 81

Walboomers JM, Jacobs MV, Manos MM  et al. Human papillomavirus is a necessary cause of invasive cervical cancer worldwide. J Pathol 1999;189,12–9

Table 1 (a) HPV positivity rates for HPV signa and comparator assaysb; (b) Concordance between HPV signaand comparator assayb

|  | (a) | RH/ INNO-LiPAb |  |  |  | HPV sign |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | 95% C.I. |  |  |  | 95% C.I. |  |
|  | Sample (n) | n | % Positive | Lower | Upper | n | % Positive | Lower | Upper |
|  | oropharyngeal squamous cell carcinomas (21) | 17 | 81% | 60% | 92% | 17 | 81% | 60% | 92% |
|  | cervical intraepithelial neoplasia 2 (47) | 47 | 100% | 92% | 100% | 38 | 81% | 68% | 90% |
|  | cervical intraepithelial neoplasia 3 (32) | 32 | 100% | 89% | 100% | 25 | 78% | 61% | 89% |
|  | All Cervical (79) | 79 | 100% | 95% | 100% | 63 | 80% | 70% | 87% |

|  | (b) | Full |  |  |  | Partial |  |  |  | Non-concordance |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  | 95% C.I. |  |  |  | 95% C.I. |  |  |  | 95% C.I. |  |
|  | Sample (n) | n | % | Lower | Upper | n | % | Lower | Upper | n | % | Lower | Upper |
| Concordance | oropharyngeal squamous cell carcinomas (21) | 17 | 81% | 60% | 92% | 4 | 19% | 8% | 40% | 0 | 0% | 0% | 16% |
|  | cervical intraepithelial neoplasia 2 (47) | 18 | 38% | 26% | 52% | 17 | 38% | 26% | 53% | 12c | 26% | 15% | 40% |
|  | cervical intraepithelial neoplasia 3 (32) | 12 | 38% | 23% | 55% | 11 | 33% | 20% | 52% | 9d | 28% | 16% | 45% |
|  | All Cervical (79) | 30 | 38% | 28% | 49% | 28 | 35% | 26% | 47% | 21 | 27% | 18% | 37% |

aHPV sign data incorporates results from follow up testing at 1/10 dilution

bRH for  oropharyngeal squamous cell carcinomas , INNO-LiPA for cervical samples

c9 +/- discordances; and 3 +/+ typing discordances

d7 +/- discordances and  2 +/+ typing discordances

Table 3 Occurrence of HPV types in  oropharyngeal squamous cell carcinomas and cervical intraepithelial neoplasia 2-3 cervical samples as detected by RH, INNO-LiPA and HPV sign

|  | Cervical (n = 79) |  | oropharyngeal squamous cell carcinomas (n = 21) |  |
| --- | --- | --- | --- | --- |
| HPV Type | RH | HPVsign | INNO-LiPA | HPVsign |
| 16 | 32 | 35 | 12 | 12 |
| 18 | 16 | 14 | 1 | 1 |
| 33 | 14 | 11 | 2 | 2 |
| 31 | 13 | 2 | 0 | 0 |
| 45 | 5 | 3 | 0 | 0 |
| 56 | 6 | 1 | 0 | 0 |
| 51 | 6 | 0 | 0 | 0 |
| 52 | 5 | 0 | 1 | 0 |
| 11 | 1 | 1 | 2 | 2 |
| 39 | 4 | 0 | 1 | 0 |
| 66 | 4 | 0 | 0 | 0 |
| 35 | 3 | 2 | 0 | 0 |
| 59 | 3 | 0 | 0 | 0 |
| 73 | 2 | 2 | 0 | 0 |
| 6 | 1 | 0 | 2 | 0 |
| 26 | 2 | 0 | 0 | 0 |
| 58 | 2 | 0 | 0 | 0 |
| 82 | 2 | 0 | 0 | 0 |
| 70 | 1 | 1 | 0 | 0 |
| 42 | 1 | 0 | 0 | 0 |
| 53 | 1 | 0 | 0 | 0 |
| 54 | 0 | 0 | 1 | 0 |
| 64 | 0 | 0 | 1 | 0 |

Table 2 Identification of High-risk HPV types

| Cervical Samples |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | RH |  |  | HPV sign |  |  | RH/HPV sign |  |  |  |  |  |
| HPV type | Mono | mixed | total | Mono | mixed | total | +/+ | +/- | -/+ | kappa | Kappa 95% CI low | Kappa 95% CI High |
| 16 | 13 | 19 | 32 | 28 | 7 | 35 | 29 | 3 | 6 | 0.74 | 0.59 | 0.89 |
| 18 | 8 | 8 | 16 | 8 | 6 | 14 | 12 | 4 | 2 | 0.71 | 0.50 | 0.91 |
| 31 | 6 | 7 | 13 | 1 | 1 | 2 | 2 | 11 | 0 | 0.76 | 0.56 | 0.96 |
| 33 | 8 | 6 | 14 | 8 | 3 | 11 | 11 | 3 | 0 | 0.86 | 0.70 | 1.00 |
| 45 | 2 | 3 | 5 | 1 | 2 | 3 | 2 | 3 | 1 | 0.48 | -0.03 | 0.98 |
| HRa | 47 | 26 | 73 | 50 | 7 | 57 | 57 | 16 | 0 | 0.35 | 0.07 | 0.64 |
| All HPV | 51 | 28 | 79 | 54 | 9 | 63 | 63 | 16 | 0 |  |  |  |

ai.e. HPV types 16, 18, 31, 33, 35, 39, 45, 51, 52, 56, 58, 59 and 68.

| Supplementary Table 1 Genotypes detected in cervical intraepithelial neoplasia 2 and cervical intraepithelial neoplasia 3 cervical samples by RH & HPV sign (n=79) |  |  |
| --- | --- | --- |
| n | RH | HPV sign |
| 1 | 6 | neg |
| 12 | 16 | 16 |
| 1 | 16 | neg |
| 6 | 18 | 18 |
| 2 | 18 | 16, 18 |
| 1 | 31 | 31 |
| 5 | 31 | neg |
| 6 | 33 | 33 |
| 1 | 33 | 16, 33 |
| 1 | 33 | 33 |
| 2 | 35 | 35 |
| 1 | 45 | 16, 45 |
| 1 | 45 | neg |
| 1 | 51 | 73 |
| 1 | 51 | neg |
| 2 | 52 | neg |
| 1 | 56 | 56 |
| 1 | 56 | 16 |
| 1 | 58 | neg |
| 1 | 59 | neg |
| 1 | 66 | neg |
| 1 | 82 | neg |
| 1 | 16, 18, 31 | 16, 18 |
| 1 | 16, 18, 31 | 16 |
| 1 | 16, 18, 31, 39 | 16, 31 |
| 1 | 16, 18, 31, 52 | 16 |
| 1 | 16, 18, 39 | 16, 18 |
| 1 | 16, 26 | 16 |
| 1 | 16, 26, 31 | 16 |
| 1 | 16, 31, 33 | 33 |
| 1 | 16, 33, 45 | 18, 33 |
| 1 | 16, 33, 52, 59 | 16 |
| 1 | 16, 35 | 16 |
| 2 | 16, 39 | 16 |
| 1 | 16, 45 | 16 |
| 1 | 16, 51, 56, 66 | 16 |
| 1 | 16, 56 | 16 |
| 1 | 16, 58 | 16 |
| 1 | 16, 66 | 16 |
| 1 | 16, 73 | 16 |
| 1 | 18, 31, 59, 66 | 18 |
| 1 | 18, 33, 56 | 18, 33 |
| 1 | 18, 51 | neg |
| 1 | 26, 70 | 70 |
| 1 | 42, 51 | 45 |
| 1 | 45, 59 | 45 |
| 1 | 51, 42, 39 | 18 |
| 1 | 52, 56, 82 | 16 |
| 1 | 53, 11 | 11 |

Supplementary Table 2: HPV genotypes identified from  oropharyngeal squamous cell carcinoma  samples by INNO-LiPA & HPV sign

| Genotypes detected |  |  |
| --- | --- | --- |
| n | INNO-LiPA | HPV sign |
| 11 | 16 | 16 |
| 1 | 11 | 11 |
| 1 | 33 | 33 |
| 1 | 6, 11 | 11 |
| 1 | 18, 39 | 18 |
| 1 | 33, 52, 54 | 33 |
| 1 | 6, 16, 64 | 16 |
| 4 | neg | neg |